## Slide 1
A2LA C63 SC6
Status Update
Megan McConnell
Spring
2024

## Slide 2: Standards
ANSI C63.4 – 337 (314)
ANSI C63.4a – 68 (63) (footnote only)
ANSI C63.5 – 17 (17)
ANSI C63.10 – 257 (250)
ANSI C63.19 – 22 (23)
ANSI C63.26 – 190 (181)
Total for C63 – 347 (324)

## Slide 3: Top Nonconforming Issues to C63 Standards
Antenna Calibration
Missing Sample Calculations
Measurement Uncertainty (Table)
Site Validation (Absorbers)
LISNs and Extension Cords
Test Setup
Notes: Full list of C63 deficiencies provided in excel format

## Slide 4: Top Nonconformances
#1 with 123 nonconformances
6.4.6 Measuring equipment shall be calibrated when:
  the measurement accuracy or measurement uncertainty affects the validity of the reported results, and/or
  calibration of the equipment is required to establish the metrological traceability of the reported results

## Slide 5: Top Nonconformances
#2 with 73 nonconformances
7.7.1 The laboratory shall have a procedure for monitoring the validity of results. The resulting data shall be recorded in such a way that trends are detectable and, where practicable, statistical techniques shall be applied to review the results. This monitoring shall be planned and reviewed and shall include, where appropriate, but not be limited to:
In the past year, the following subclauses were cited:
  a, c, d, e, i, j
Notes: Full list of C63 deficiencies provided in excel format

## Slide 6: Top Nonconformances
#3 with 59 nonconformances
7.5.1 The laboratory shall ensure that technical records for each laboratory activity contain the results, report and sufficient information to facilitate, if possible, identification of factors affecting the measurement result and its associated measurement uncertainty and enable the repetition of the laboratory activity under conditions as close as possible to the original. The technical records shall include the date and the identity of personnel responsible for each laboratory activity and for checking data and results. Original observations, data and calculations shall be recorded at the time they are made and shall be identifiable with the specific task.

## Slide 7: Regulator Activity
FCC – 230 (230) Recognized – 75 US Test Labs – 19 (17) TCBs
  2 Year Cycle/Extensions
  Lack of Definitions/rules
  Timely Publications
  New R&O and FPRM
ISED – 240 (242) Recognized – 64 US
  Limitations on Scopes/Checklists
  CISPR 16-1-4
ASCA – 59 (78) Recognized (Probably?)
  80 Scope on A2LA Directory

## Slide 8
A2LA 5202 Presidents Court
Suite 220
Frederick, MD 21703
301 644 3248 Main
301 662 2974 Fax
info@A2LA.org
www.A2LA.org
 2023 by A2LA
All rights reserved. No part of this document may be reproduced in any form or by any means without the prior written permission of A2LA.